## Slide 1: Running a (magnetic) measurement laboratory – the value-shop model, uncertainty and calibration, ignorance, avatars and twins
Stephan Russenschuck, Gorini 2020

## Slide 2: Iron Dominated Magnets

## Slide 3: Coil Dominated Magnets
B = 4 T Bs = 3.69 T
B = 8.33 T Bs = 7.77 T

## Slide 4: Class 1 and 2 Superconducting Magnets
Class 1
Class 2

## Slide 5: Accelerator Magnets: A Multiphysics Problem
Material science: Superconducting cable, steel, insulation
Mechanics and large-scale mechanical engineering
Vacuum technology
Cryogenics (superfluid helium)
Metrology and alignment
Electrical engineering (power supplies, leads, quench detection and magnet protection
Magnetic field measurements (MM)
Analytical and numerical field computation

## Slide 6: MM as a Multidisciplinary Task
Beam physics (understanding the user requirements)
Mechanics (support structures, translation stages, vibration control)
Electronics for signal processing (DAQ, filtering, amplification, integration)
Architecture (laboratory space)
Electrical engineering (power converters)
Statistics (uncertainty analysis, robust design)
Potential theory (mathematical physics) for data postprocessing
Programming
Project management (procurement of specialized components, assembly, metrological characterization)
Line management (multi-cultural, multi-disciplinary, multi-level, limited duration)

## Slide 7: Environment: We Have Come a Long Way
ISR-Tunnel Segment: The most stable environment at CERN

## Slide 8: Architecture: Integration Studies

## Slide 9: The New Measurement Laboratory 311
Air-condition system (1.4 + 0.4 MCHF)
Avoid stratification (laser applications)
Air speed < 0.4 m/sec (laminar)
+- 2 oC in large hall, +- 0.4 oC in “measuring room”
Lighting:
1400 Lumen
5500 K (warm daylight)
Power converter cluster,
motorized switches,
PLC interlocked distribution
feedboxes
Intrinsic safety

## Slide 10: The Measuring Room: A Lesson Learnt
Vibration control:
locally
Thermal control:
globally

## Slide 11
Hirsch-index R&D: New measurement technologies and devices (not specific to a project). Theory and proof of principle. Operational (R&D) budget. TECH/DOCT students, co-supervision with universities.
Project-driven R&D: New bench to be developed for specific measurement tasks. Project budget (CERN accelerators and approved physics experiments).
Project-driven development: Preparation of new components (shafts, coils, FFMM class). Project budget.
Win-win collaborations (partner institutes, industry): Development of special purpose equipment. External funding sources (material and personnel).
Industrial-type services: Standard measurements with existing equipment. Operational, project budgets or external funding. Technology transfer (licencing) to partner institutes or industry.
Modus Operandi of the MM Section

## Slide 12: The Measurement Paradigm
Measurements for design
  Material properties and design choices
  Software and model validation
  Dynamic effects
Measurements for magnet production
  Prototype qualification
  Quality assurance (error tracking); acceptance testing
  Magnet to magnet reproducibility (check tolerances)
  Shimming
  Requalification after repair
Measurements for accelerator operation
  Field description (feed forward compensation)
  Field quality (resonances)
  Online monitoring (pre-cycling)

## Slide 13: Beam Physicists, Magnet Designers, and Measurement Engineers
Everybody believes in measurements, but the measurement engineer himself.
Nobody believes in field simulations, but the field computation expert himself.

## Slide 14: C3 Coherence
Machine
operation
Beam
optics
Magnet
design
Magnet
manufacture
Magnetic
measurements
Measurement artifacts,
Unexpected physical phenomena,
Modeling and manufacturing errors
Alignment tolerances
Dynamic effects
Empirical Magnet
Model
Online monitoring for machine control
Magnet
requirements
Magnet spec.
Measurement needs

## Slide 15: A Few Years Later

## Slide 16: Avatars and Twins
2D Field map Good field region
3D Field map
Generalized field description from measurements (and simulations) with post-processing steps based on Kirchhoff’s theorem

## Slide 17: Maxwell’s House

## Slide 18: The Avatar
Numerical model (FEM/BEM
Eigenfunctions)
Magnetic flux
density in W
Predicted signal
(Voltages, flux increments)
Magnetic
Measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density)
Collocation, Least-squares, Bayesian inversion (prior)

## Slide 19: Special Case: Multipoles (Solving of Boundary Value Problems)
2. Calculate the field components
Eigensolutions of the Laplace equation
(coefficients are not known)

## Slide 20: Solving of Boundary Value Problems
3. Measure the field on the domain boundary and perform Fourier
analysis (develop into the eigenfunctions). Coefficients known here.
4. Compare the known and unknown coefficients
5. Put this into the original solution for the entire air domain

## Slide 21: Rotating Coil Systems

## Slide 22: Magnetic Measurement Techniques
Fixed/rotating coils
Hall sensor
FMR
Stretched wires
NMR
Sensor field range [T]
relative uncertainty [-]
Fixed coils
Hall sensor
Rotating coils
FMR
Stretched wires
Longitudinal sensor size (mm)
NMR
Transversal sensor size (mm)

## Slide 23: Magnetic Measurement Techniques
Rotating coil systems
3D Hall mapper
Stretched-wire
Slow, least accurate
Stationary induction-coil
magnetometers
Expensive
Integrated fields only
Fast-ramped magnets only
Definition of the space-curve
Good for magnet-to-magnet reproducibility

## Slide 24
Asset management
  Inventory
  Calibration data and history
  Preventive maintenance
Operations management
  Magnetic measurement requests
  Schedule and equipment allocation
  Preparation of measurement bench and MM execution
Project management
  Design (reviews) and procurement (tendering, production follow up)
Business intelligence
  Reporting
  Traceability
Resource planning
  Material budget and resources (personnel)
Management Tasks

## Slide 25
Scheduling (Kanban Board)
ATLASSIAN Jira

## Slide 26
Space Management
Booking tool (Midas)
Bench status display
(Drupal 8 Website)

## Slide 27
MM requests
Kanban Board
NORMA
MTF
EDMS
MM Equipment
Calibration data
MM report
FSU job request
Operations management
Database “management”
Asset management
Project management and
Business intelligence
Released data
FFMM config.
Raw data
Processed data
Material
Kn
Traveller
NC magnet
SC magnet
Bench preparation
Client
Client
The Problem Setting

## Slide 28
Certification Requirements (ISO/IEC 17025)
General
Accountability, impartiality, confidentiality
Resources
Staffing, training
Environment, equipment, traceability
Processes
Customer relation
Proper naming convention
Electronic logbooks
Propagation of uncertainty (GUM)
Record equipment meta data
Restrict access to calibration equipment
Automatic generation of protocols
Centralized data storage
Management systems
Quality assurance and audits
Production readiness
Documentation

## Slide 29: The MM Value-Shop Model

## Slide 30: Measurement Dataflow

## Slide 31
EAM (Enterprise Asset Management)
Why InforEAM
Used throughout CERN
service, maintenance and support provided by a dedicated Section at CERN
Integration with other CERN systems (EDMS, MTF, TREC, etc.)
Can be integrated with other web-services
Customisable with some restrictions

## Slide 32
Magnetic Measurement Requests and Workorders
Equipment code
Requestor’s timeline
Other information from requestor
Information from NORMA / MTF
Assignment of resources
Scheduling
Raw and processed data
Released
(MM Protocol)
Synopsis
(MM Report)
Request
Phase
Preparation
Phase
Execution
Phase
Post-processing
Phase
Work-order
Equipment used
MM procedures and parameters
Definition of MM procedure
Development and procurement

## Slide 33
Device Registration and Asset Codes
  Crucial for magnetic measurement
  Holds/pushes data required in FFMM
  Requires regular maintenance and calibration
  Exceeds a defined financial value
  Unique and frequently used

## Slide 34
The Interfaces
MM Request
(Client)
Workorders (Events in EAM)

## Slide 35: Naming Conventions
Magnetic field-sensing devices
  Scalar fields: Nuclear magnetic resonance (NMR), electron spin resonance (ESR), superconducting quantum interference devices (SQUID
  Vector fields: micro-electromechanical, magnetoresistance, Hall effect, fluxgate.
  Integral (vector)-field quantities: Vibrating sample, fibre-optical (Faraday effect), torque (cantilever), induction coil, and single-stretched wire.
Transducer
  Comprises the sensor and any axillary electronics needed to convert the magnetic flux (density) into an electric voltage at its connection terminals.
Induction-coil devices
  Rotating-coil magnetometers, translating-coil magnetometers, stationary induction-coil magnetometers (fluxmeters)

## Slide 36: Naming Convention
A single coil of the magnetometers is called an induction coil.
  An ordered set of induction coils is called an array. A cylindrical induction-coil array is called a segment.
    A shaft is made from one or more segments.
      A chain is a series of connected shafts.
A rotating-coil scanner, and a translating-coil scanner are devices consisting of a shaft or induction coil-array mounted on a one-dimensional displacement stage.
A displacement stage for at least two dimensions is called a mapper.
The measurement system comprises the bench, the actuators, motor drives, the transducers, the data acquisition system, the digital integrator, and the software framework (FFMM).
  A bench is the mechanical support structure for the magnet and the measurement system.
    A workspace comprises the measurement system, the magnet, and the allocated floor space in the measurement lab.

## Slide 37: Polarity Convention, Positioning, and Alignment
Consistently
oriented
Non-consistent inner and outer
orientation: to be avoided
Definition of d and q axes
and the (mis)-alignment angles

## Slide 38
Asset Codes (Coil Magnetometers)
CRMMMSH_AA-000000657
CRMMMSE_AA-00000268
CRMMMSE_AA-00000245
CRMMMSE_AA-00000233
CRMMMCO_AA-00004450
CRMMMCO_AA-00001949
CRMMMCO_AA-00004550
CRMMMCO_AA-00000583
CRMMMCO_AA-00089001
CRMMMCO_AA-00002229
CRMMMCO_AA-00005768
CRMMMCO_AA-00006778
CRMMMCO_AA-00000230
Shaft asset code
Segment asset code
Shaft
Total number of coils
Total geometric length
Plane position (mm)
Geometric radius
Tangential/Radial
Connector position
As link in a chain (y/n)
Calibration validity
Material orders
Drawings
Segment
Material type
Calibration date
Kn (generic)
Kn* (calibrated)
Calibration raw data {N}
Drawings
Cabling scheme
Coil
Track dimensions
Contraction factor
Material type (support)
Material type (wire)
Number of turns
Consistent orientation (y/n)
Coil inner length
Coil inner width
Winding width (trans)
Winding width (longitudinal)
Coil height
Number of layers
Coil surface (theoretical)
Coil surface (calibrated)
Resistance
Temperature
Calibration date
Material orders
Calibration raw data
Drawings
Design files
Coil in segment
Coil asset code
Coil position
Radius
Beta
Alpha
Delta alpha {N}
Delta beta {N}
Normal orientation (y/n)
Polarity (+/-)
Connector type
Pin number A term
Pin number B term
Segment in shaft
Segment position

## Slide 39
MM Equipment Lifecycle Measurements
Design
Procurement
Assembly
Metrological
characterization
Calibration
Preventive
maintenance
Corrective
maintenance
Decom-
missioning
Assign
Equipment
Code
Operation
Variant
N+1

## Slide 40: Data Flow
Meta data are equipment codes, calibration data, environmental data, logfiles
Raw (analog) signals coming from the transducers (induction coils, Hall sensors), which are the voltages over their connection terminals.
Raw (digital) data are the signals at the port of an AD converter and of a synchronized timing card which allows to establish the numerical values of U(t). The two functionalities are combined in the digital integrator card (FDI) developed at CERN, which yields the measured quantity F(t).
Pre-processed data are raw data after some post-processing, which include averaging, filtering, integrator-drift correction, and correction for the instruments input resistance, among others.
Output data (MM results) are pre-processed data that has undergone additional processing steps, such as digital bucking, deconvolution, and Fourier transform. The results are stored in measurement protocols that are automatically uploaded to the database by means of the software framework (FFMM) and the asset management system (Infor EAM) .
Released data are validated and analyzed data that are compared to the magnet specification (conform, nonconform, uniformity). Released data are presented in reports uploaded on the database via CERN's Engineering Data Management System (EDMS).

## Slide 41
Measurement script
Parameters
Raw data
Pre-processed data
Output data
Protocols
Released data
Reports
Data Flow
Bench
setup
DAQ
Data
processing
Analysis

## Slide 42: Uncertainty and Mitigation Measures
Application
MM-Method
Types of uncertainty
Mitigation
Long (SC) magnets
(LHC arc dipoles
and quads)
Large aperture quads
(HL-LHC inner triplets,
SuperFRS multiplets)
Spectrometers
Solenoids
(Anti-proton traps)
Strongly curved
magnets
Long rotating
coil magnetometers
(integrated mulitpoles)
Stretched wire
(magnetic axis)
Hall mapper
(field maps for
track reconstruction)
Moving fluxmeter
Short (isoperim.) coils
Solenoidal-field
transducer
Curved wire
Approximation
(incomplete basis)
Calibration
Systematic
Random
Oversampling
Integration
Bucking
Cross-calibration
In-situ calibration
Intrinsic
(false assumptions)
Ignorance
(incomplete
physical model)
Bayesian inference
Deconvolution
Local Maxwellification
Sensor fusion
Experimental magnets
Gantry
Cauchy estimate

## Slide 43: Sources of Uncertainty
Random errors (capture and read-out noise). Oversampling (repeated measurement, averaging, learnt statistical analysis (type A evaluation)), signal compensation (bucking), a stable mechanical support, and vibration damping. These errors are the least concern in magnetic measurements.
Systematic errors. Calibration, ideally traceable to primary metrological standards. Tight control on the environment (laboratory conditions), polarity inversion, sensor reversal, go-and-return (compensate for integrator drift), comparison with references (type B evaluation).
Correlation (covariance), propagation of uncertainty and design of experiments: Guide to the Expression of Uncertainty in Measurement (GUM)

## Slide 44: Sources of Uncertainty
Operator (gross) errors may arise from using the wrong calibration data for an instrument, or attributing the measurements to a wrong equipment code. Linking asset management with DAQ systems.
Approximation errors from the projection of the measured data onto Fourier polynomials (of order N), truncated Taylor series, wavelet expansions, or sparse finite-element meshes. Convergence studies.
Intrinsic errors from a (false) assumption of linearity of the sensor response or by non-orthogonal sensors. “Maxwellification” in a post-processing step (avatars).
Ignorance from model reduction. Limitation of the electromagnetic subsystem, Faraday paradox, neglecting mechanical effects (magnetostriction) or thermal effects in permanent magnets. In a number of cases including three-dimensional eddy-current and vector hysteresis effects, no avatar in the required 10-4 range can be constructed.
  Machine learning techniques (twins).

## Slide 45: Use of Faraday’s Law
Flux rule
U = dF/dt
Auchmann, Kurz, Russenschuck:
A note on Faraday paradoxes

## Slide 46: Mitigation Measures (Intrinsic Errors and Ignorance)
Be careful and check the validity of apparently “well-known” equations. The Maxwell house is not valid for switches, sliprings, multi-connected domains, non-orientable manifolds.
Be particularly careful when you use data from different transducers (in physics you cannot add apples and oranges, on the computer you can). Regularization techniques (avatars) can help.
No 10-4 predictive models exist for hysteresis (iron, permanent magnets), 3D eddy currents, magnetostriction (multi-physics) Therefore no avatar can be created. Twins are a combination of numerical models and measurements, with machine learning and empirical rules.

## Slide 47: Mitigation Measures (Systematic and Gross Errors)
Calibration: to establish the transfer function of a given instrument by comparison with a reference instrument of known accuracy
  cross-checks with instruments of comparable uncertainty to gain confidence or eliminate gross errors
  in-situ calibration: use the magnet to be tested as the reference via a previous measurement with a reference instrument
Traceability: an unbroken chain of comparisons, each with a stated uncertainty, from a measurement to a primary standard.
  Fundamental concept to certify a measurement; maintaining systematic records, databases, documented procedures

## Slide 48: CERN Dipole Field Reference
Old CERN ISR bending dipole, cycled up to 1 T in a rigorously reproducible way
Yearly mapping to establish field profile, averages and reproducibility
We map By, but since the field is not perfectly uniform we have also Bx, Bz  the resulting (small) error is propagated to all instruments calibrated in this reference !
high precision NMR probe used to map the field

## Slide 49
3D-Helmholtz-Coil

## Slide 50: Mitigation Measures (Systematic and Gross Errors)
Flexible C++ Framework for Magnetic Measurements
Web-services
(XML (Soap))
Text-file
SQL

## Slide 51
Use Case: Quadrupole-Shaft Calibration
R/W
Notes: Quadrupole calibration – Shaft mounted inside the reference magnet

## Slide 52: FFMM Roles
Measurement
software
framework
Measurement
software
application
Device 1
Device n
Test engineer context
Application Operator context
...
Developper context

## Slide 53: FFMM Current State of Implementation

## Slide 54: FFMM Onboard Analysis
Starting point: Post-processing of raw data in Matlab
Provide features to run analysis online during data taking
    Immediate feedback
    Generation of protocols
Multiple approaches
    a) Use Matlab generated code for analysis (compiled C++)
      Code sharing between offline post-processing scripts and online analysis, quick implementation
    b) Use native C++ implementation
      Use library for common needs (linear algebra, solvers, FFT calculation), no dependencies on Matlab

## Slide 55
Combination of InforEAM, EDMS, and custom Oracle database
Trade-off between ease of access and ease of implementation
    Meta-data into InforEAM
    Results into Oracle database
    Raw data files and measurement parameters into EDMS
Storage of all measurement parameters (Links in Inform EAM)
    Devices
    Settings
    Calibration data
No raw data are discarded → reconstruction possible
    In case of incorrect analysis parameters or errors
    Future investigation of measurement results
FFMM Database Management

## Slide 56
Certification Requirements (ISO/IEC 17025)
General
Accountability, impartiality, confidentiality
Resources
Staffing, training
Environment, equipment, traceability
Processes
Customer relation
Proper naming convention
Electronic logbooks
Propagation of uncertainty (GUM)
Record equipment meta data
Restrict access to calibration equipment
Automatic generation of protocols
Centralized data storage
Management systems
Quality assurance and audits
Production readiness
Documentation
MERIT, Budget
TECH, DOCT, Fellow, Staff
InforEAM
C3
Equipment codes
FFMM, InforEAM
Robust design
FFMM, InforEAM
CERN access, Jira, Midas
FFMM
FFMM - InforEAM – Oracle DB
QA manager
Value-shop model
EDMS, Publications